	
	 
	 
	

	PROCEDURE
	

	VITROS Chemistry Products Calibrator Kit 1
	Calibrator Kit 1

	 
	[image: image2.png]



	188 2208


	
	

	
	PROCEDURE

	Calibrator Kit 1
	


	
	

	PROCEDURE
	

	
	Calibrator Kit 1




Rx ONLY

 
Intended Use
For in vitro diagnostic use only.
VITROS Chemistry Products Calibrator Kit 1 is used to calibrate VITROS 250/350/5,1 FS/4600/XT 3400 Chemistry Systems and VITROS 5600/XT 7600 Integrated Systems for the quantitative measurement of BUN/UREA, Ca, CREA, GLU, LAC, Li, Mg, PHOS, SALI, THEO, and URIC.
 
Reagents
These calibrators are prepared from processed bovine serum to which organic analytes, electrolytes, stabilizers, and preservatives have been added.
The companion diluents are prepared from processed water to which inorganic salts have been added.
 
Nominal Values and Traceability
Nominal values are representative target concentrations used during the calibrator manufacturing process. The particular Supplementary Assigned Value (SAV) for an analyte in each vial is the generation-specific assigned concentration for VITROS Chemistry Products Slides, and is provided on the calibration media. To view this value, refer to the operating instructions for your system. Refer to the analyte specific Instructions for Use for additional calibration information.
 
Nominal Values
 
	Analyte
	
	Calibrator Vial 1 * 
	Calibrator Vial 2 * 
	Calibrator Vial 3 * 
	Units
	Calibrator Vial 1 * 
	Calibrator Vial 2 * 
	Calibrator Vial 3 * 
	Units

	Urea
	(BUN/UREA)
	4
	50
	115
	mg/dL
	1.4
	17.8
	41.0
	mmol/L

	Calcium
	(Ca)
	13.8
	9.3
	1.8
	mg/dL
	3.44
	2.32
	0.45
	mmol/L

	Creatinine
	(CREA)
	0.5
	1.5
	13.2
	mg/dL
	44
	133
	1167
	µmol/L

	Glucose
	(GLU)
	35
	300
	600
	mg/dL
	1.9
	16.7
	33.3
	mmol/L

	Lactate
	(LAC)
	0.5
	3.0
	12.0
	mmol/L
	0.5
	3.0
	12.0
	mmol/L

	Lithium
	(Li)
	0.2
	1.5
	4.0
	mmol/L
	0.2
	1.5
	4.0
	mmol/L

	Magnesium
	(Mg)
	0.6
	3.8
	9.5
	mg/dL
	0.2
	1.6
	3.9
	mmol/L

	Phosphorus
	(PHOS)
	1.0
	5.4
	12.6
	mg/dL
	0.32
	1.74
	4.07
	mmol/L

	Salicylate
	(SALI)
	1
	23
	38
	mg/dL
	0.1
	1.7
	2.8
	mmol/L

	Theophylline
	(THEO)
	1.5
	19.0
	36.0
	µg/mL
	8.3
	105.5
	199.8
	µmol/L

	Uric Acid
	(URIC)
	0.7
	6.0
	16.5
	mg/dL
	42
	357
	981
	µmol/L


* Concentration after calibrator is reconstituted with 3.0 mL of corresponding diluent.
 
Traceability of Values
 
	Analyte/Chemistry
	
	Reference Material
	Comparative Method

	Urea
	(BUN/UREA)
	NIST SRM 912
	CDC (Urease/GLDH) 1 

	Calcium
	(Ca)
	NIST SRM 915
	Atomic Absorption 2 

	Creatinine
	(CREA)
	NIST SRM 914
	HPLC 3  

	Creatinine
	(CREA)
	NIST SRM 914
	IDMS 4 

	Glucose
	(GLU)
	NIST SRM 917
	AACC/CDC (Hexokinase/G6PDH) 5 

	Lactate
	(LAC)
	USP certified sodium lactate
	HPLC 6 

	Lithium
	(Li)
	NIST SRM 924
	Flame Atomic Absorption 7 

	Magnesium
	(Mg)
	NIST SRM 929
	Flame Atomic Absorption 8 

	Phosphorus
	(PHOS)
	NIST SRM 200
	Phosphomolybdate/p-semidine HCI 9 

	Salicylate
	(SALI)
	USP certified sodium salicylate
	HPLC 10 

	Theophylline
	(THEO)
	USP certified theophylline
	HPLC 11 

	Uric Acid
	(URIC)
	NIST SRM 913
	Uricase/CDC 12 


 
Warnings and Precautions
For in vitro diagnostic use only.
 
	WARNING:
	While these products are bovine in origin, they should be handled using the same precautions as with any other blood or blood-derived product.


 
	WARNING:
	The packaging (vial stopper) of this product contains dry natural rubber, which may cause allergic reactions in some individuals.


Not all products and systems are available in all countries.
 
Reconstitution
 
	Caution:
	Do not use visibly damaged product or product with incompletely sealed packaging.


 
	Note:
	Each bottle of calibrator lyophilate has a corresponding diluent labeled by number. Use the appropriate diluent in the reconstitution of the lyophilate.


1. Materials should be at room temperature, 18–28 °C (64–82 °F), before reconstitution. Vials should sit out at room temperature approximately 60 minutes when taken from freezer storage.
2. Slowly invert the diluent bottle several times to mix the contents thoroughly. DO NOT SHAKE.
3. Gently tap the lyophilate vial on the counter several times to dislodge any material adhering to the stopper.
4. Remove the seal and stopper from each bottle just before adding the diluent. Do not leave vials unstoppered.
5. Add 3.0 mL of the appropriate diluent to each vial. Use a clean, dry pipette for each vial. A Class A volumetric pipette or an automated pipette of equivalent accuracy is recommended because of the importance of this reconstitution procedure to the accuracy of the results. Discard any remaining diluent.
6. Replace the stopper and hold it firmly in place. Invert the vial gently. DO NOT SHAKE.
7. Reconstitution, with occasional inversion, may take up to 30 minutes. Visually verify that all freeze-dried material is dissolved prior to use.
8. Keep all fluids tightly stoppered when not in use. At the time of reconstitution, it is recommended the operator date and initial the vial.
9. Reconstituted product should be used immediately or stored in the refrigerator between 2–8 °C (36–46 °F).
 
Storage
Storage and Stability
 
	Reagent
	Storage Condition
	Stability

	Unopened
	Frozen
	≤-18 °C (≤0 °F)
	Until expiration date

	Reconstituted
	Refrigerated
	2–8 °C (36–46 °F)
	≤ 24 hours if tightly stoppered


Refer to the analyte specific Instructions for Use for special calibration precautions.
 
Materials Provided
VITROS Chemistry Products Calibrator Kit 1
· 4 vials each of lyophilized calibrator 1, 2, and 3
· 4 vials each of calibrator diluent 1, 2, and 3
· 12 calibrator bar code labels (For use on VITROS 4600/XT 3400 Chemistry System and 5600/XT 7600 Integrated Systems)
 
Materials Required but Not Provided
A Class A volumetric pipette or an automated pipette of equivalent accuracy for addition of diluent to lyophilate
 
Testing Procedure
 
	Note:
	Be sure to use components from the same kit lot number.


1. If necessary, remove reconstituted material stored in the refrigerator.
2. Mix vial thoroughly by gently inverting several times. DO NOT SHAKE.
3. Place fluid in a cup and cover with a pierceable cap.
4. Restopper the vial and immediately return it to the refrigerator.
5. Bring the cup to room temperature, 18‍–‍28 °C (64–82 °F), before analysis (approximately 10 minutes for refrigerated material).
6. Analyze according to operating instructions for your system.
7. Discard any unused portion in the cup following testing.
8. Use refrigerated fluid within 24 hours.
 
Limitations
The commutability of the VITROS Chemistry Products Calibrator Kit 1 for BUN/UREA, Ca, CREA, GLU, LAC, Li, Mg, PHOS, SALI, THEO, and URIC has been demonstrated among VITROS MicroSlide methods. Commutability of this calibrator has not been established with other methods.
 
References
1. Sampson EJ, et al. A coupled-enzyme equilibrium method for measuring urea in serum: optimization and evaluation of the AACC study group on urea candidate reference method. Clin. Chem. 26:816-26; 1980.
2. Cali JP, et al. Atomic Absorption. NBS Reference Method (modified). Clin. Chem. 19:1208; 1987.
3. Ambrose RT, Ketchum DF, Smith JW. Creatinine Determined by “High Performance” Liquid Chromatography. Clin. Chem. 29: 256-259; 1983.
4. Siekmann L. Measurement of creatinine in human serum by isotope dilution mass spectrometry. J Clin Chem Clin Biochem  23:137–144; 1985.
5. Neese JW, Duncan P, Bayse DD, et al. Development and evaluation of a hexokinase/glucose-6-phosphate dehydrogenase procedure for use as a national glucose reference method. HEW Publication No. (CDC) 77-8330. HEW. USPHS, Centers for Disease Control, 1976.
6. Smith JW, Ambrose RT. Determination of Lactic Acid in Human Serum by Ion Exchange Chromatography. Internal Eastman Kodak Company Report. 1982.
7. A Reference Method of Determination of Lithium in Serum. U.S. Dept. of Comm., NBS Publication 260-69.
8. Kaplan L, Pesce A. Clinical Chemistry: Theory, Analysis, and Correlation. CV Mosby; 1069; 1984.
9. Dryer RL, Tamnes AR, Routh JL. The Determination of Phophorous and Phosphatase with N-Phenyl-p‑phenylenediamine. J. Biol. Chem. 222:177; 1957.
10. Cham BE, Johns D, Bochner F, Imhoff D, Rowland M. Simultaneous Liquid-Chromatographic Quantitation of Salicylic Acid, Salicyluric Acid, and Gentisic Acid in Plasma. Clin. Chem. 25:1420-1425; 1979.
11. Lauff J. A Reference Procedure for the Determination of Theophylline and Related Xanthines in Serum by Dynamic Ion-Exchange HPLC. J. Chrom. 417:99-109; 1987.
12. Duncan P, Gochman N, Cooper G, et al. Development and Evaluation of a Candidate reference Method for Uric Acid in Serum. CDC, US Dept. of HHS, Atlanta, GA; 1979.
 
Glossary of Symbols
[image: image1.png]The following symbols may have been used in the labeling of this product.

Do Not Reuse

Use by or Expiration Date
(Year-Month-Day)

el

LOT Batch Code or Lot Number

Serial Number

Catalog Number or
Product Code

m

RO [

Caution

Keep Dry (Protect from
Moisture/Humidity)

Manufacturer

Date of Manufacture

Authorized Representative
in the European Community

Corrosive

Health Hazards

/ﬂf Upper Limit of Temperature
Lower Limit of Temperature
Temperature Limitation

Consult Instructions for Use

Attention: The Instructions for
Use (IFU) has been updated

For use in Slide Supply 1
For use in Slide Supply 2

S| Units
Conventional Units

Value

Flammable

Acute Toxicity

® S [ilipgE=~

Range

Range of Means

Midpoint

Revised

Supersedes

Contains Sufficient for

n" Tests

in vitro Diagnostic

Medical Device

Der Griine Punkt (the Green Dot).
Manufacturer follows certain

packaging material waste disposal
management regulations

Estimated within-lab SD

Serious Health Hazards

Environmental or
Aquatic Toxicity





 
Revision History
 
	Date of Revision
	Version
	Description of Technical Changes

	 
	 
	 


 
	When this Procedure is replaced, sign and date below and retain as specified by local regulations or laboratory policies, as appropriate.

	 
	 
	 

	Signature
	 
	Obsolete Date


 
 
	 
	This Procedure was derived from the following Ortho Clinical Diagnostics document: Pub. No. J11322_EN Version 13.0
	The information in this Procedure document has been provided to customers using Ortho Clinical Diagnostics products and was current as of the publication date. This document is not intended for commercial sale or distribution and may not be provided to other users without the written permission of Ortho Clinical Diagnostics. The Instructions for Use documents are the authorized source of information for VITROS Chemistry Products and VITROS Immunodiagnostic Products.


	
	Ortho Clinical Diagnostics is not responsible for any editing, line alteration or other revision, intentional or otherwise, to this information.
	



	

	Ortho Clinical Diagnostics is not responsible for any editing, line alteration or other revision, intentional or otherwise, to this information.
	


	
	Ortho Clinical Diagnostics is not responsible for any editing, line alteration or other revision, intentional or otherwise, to this information.
	




[image: image2.png]