	
	 
	 
	

	PROCEDURE
	ALT

	VITROS Chemistry Products ALT Slides
	Alanine Aminotransferase

	 
	[image: image7.png]



	165 5281


	
	

	ALT
	PROCEDURE

	Alanine Aminotransferase
	


	
	

	PROCEDURE
	ALT

	
	Alanine Aminotransferase




Rx ONLY

 
Intended Use
For in vitro diagnostic use only.
VITROS Chemistry Products ALT Slides quantitatively measure alanine aminotransferase (ALT) activity in serum and

plasma using VITROS 250/350/5,1 FS/4600/XT 3400 Chemistry Systems and the VITROS 5600/XT 7600 Integrated

Systems. 
Summary and Explanation of the Test
Alanine aminotransferase is present in high activity in liver, skeletal muscle, heart, and kidney. Serum ALT increases rapidly in liver cell necrosis, hepatitis, hepatic cirrhosis, liver tumors, obstructive jaundice, Reye’s syndrome, extensive trauma to skeletal muscle, myositis, myocarditis, and myocardial infarction. 1 
 
Principles of the Procedure
The VITROS ALT Slide is a multilayered, analytical element coated on a polyester support.
A drop of patient sample is deposited on the slide and is evenly distributed by the spreading layer to the underlying layers. The spreading layer contains the ALT substrates L-alanine and sodium α-ketoglutarate. Alanine aminotransferase catalyzes the transfer of the amino group of L-alanine to α-ketoglutarate to produce pyruvate and glutamate. Lactate dehydrogenase (LDH) then catalyzes the conversion of pyruvate and NADH to lactate and NAD+.
The rate of oxidation of NADH is monitored by reflectance spectrophotometry. The rate of change in reflection density is proportional to enzyme activity.
Test Type and Conditions
 
	Test Type
	VITROS System
	Approximate Incubation Time
	Temperature
	Wavelength
	Reaction Sample Volume

	Multiple-point rate
	5600, 4600, 5,1

FS, 250/350, XT

7600, XT 3400
	5 minutes
	37 °C (98.6 °F)
	340 nm
	11 µL


Not all products and systems are available in all countries.
 
Reaction Scheme
 
	alanine + α-ketoglutarate
	ALT
[image: image1.png]\




pyridoxal-5-phosphate
	pyruvate + glutamate

	pyruvate + NADH + H+
	LDH
[image: image2.png]\




	lactate + NAD+


	


 
Warnings and Precautions
For in vitro diagnostic use only.
 
	WARNING:
	Take care when handling materials and samples of human origin. Since no test method can offer complete assurance that infectious agents are absent, consider all clinical specimens, controls, and calibrators potentially infectious. Handle specimens, solid and liquid waste, and test components in accordance with local regulations and CLSI Guideline M29 2  or other published biohazard safety guidelines.


For specific warnings and precautions for calibrators, quality control materials, and other components, refer to the Instructions for Use for the appropriate VITROS product, or to other manufacturer’s product literature.
 
Reagents
 
	 
	Slide Diagram
	 

	Slide Ingredients
	[image: image3.png]



	1. Upper slide mount
2. Spreading layer (BaSO4)
· sodium α–ketoglutarate
· L-alanine
3. Reagent layer
· buffer, pH 8.0
· lactate dehydrogenase
· NADH
· pyridoxal-5-phosphate
4. Support Layer
5. Lower slide mount

	Reactive Ingredients per cm2
	
	

	Lactate dehydrogenase (porcine muscle) 0.12 U; L-alanine 0.86 mg; sodium α-ketoglutarate 54 µg; nicotinamide adenine dinucleotide, reduced 35 µg; and sodium pyridoxal-5-phosphate 11 µg.
	
	

	Other Ingredients
	
	

	Pigment, binders, buffer, surfactants, cross-linking agent and stabilizer.
	
	


 
Reagent Handling
 
	Caution:
	Do not use slide cartridges with damaged or incompletely sealed packaging.


· Inspect the packaging for signs of damage.
· Be careful when opening the outer packaging with a sharp instrument so as to avoid damage to the individual product packaging.
 
Reagent Preparation
 
	IMPORTANT:
	The slide cartridge must reach room temperature, 18–28 °C (64–82 °F), before it is unwrapped and loaded into the slide supply.


1. Remove the slide cartridges from storage.
2. Warm the wrapped cartridge at room temperature for 30 minutes when taken from the refrigerator or 60 minutes from the freezer.
3. Unwrap and load the cartridge into the slide supply.
 
	Note:
	Load the cartridges within 24 hours after they reach room temperature, 18–28 °C (64–82 °F).


 
Reagent Storage and Stability
VITROS ALT Slides are stable until the expiration date on the carton when they are stored and handled as specified. Do not use beyond the expiration date.
 
	Reagent
	Storage Condition
	Stability

	Unopened
	Refrigerated
	2–8 °C (36–46 °F)
	Until expiration date

	
	Frozen
	≤-18 °C (≤0 °F)
	Until expiration date

	Opened
	On-analyzer
	System turned on
	≤ 4 weeks

	
	On-analyzer
	System turned off
	≤ 2 hours


Verify performance with quality control materials:
· If the system is turned off for more than 2 hours.
· After reloading cartridges that have been removed from the slide supply and stored for later use.
 
Specimen Collection, Preparation and Storage
Specimens Recommended
· Serum
· Plasma:
–
EDTA
–
Heparin
 
	IMPORTANT:
	Certain collection devices have been reported to affect other analytes and tests. 3  Owing to the variety of specimen collection devices available, Ortho Clinical Diagnostics is unable to provide a definitive statement on the performance of its products with these devices. Confirm that your collection devices are compatible with this test.


 
 
Specimens Not Recommended
Do not use hemolyzed specimens. 4 
 
Serum and Plasma
Specimen Collection and Preparation
Collect specimens using standard laboratory procedures. 5  ,   6 
 
	Note:
	For details on minimum fill volume requirements, refer to the operating instructions for your system.


 
Patient Preparation
No special patient preparation is necessary.
Special Precautions
Centrifuge specimens and remove the serum or plasma from the cellular material within 3 days of collection. 7 
 
Specimen Handling and Storage
· Handle and store specimens in stoppered containers to avoid contamination and evaporation.
· Mix samples by gentle inversion and bring to room temperature, 18–28 °C (64–82 °F), prior to analysis.
 
	IMPORTANT:
	Do not freeze the specimen.


Specimen Storage and Stability 7 
 
	Storage
	Temperature
	Stability

	Room temperature
	18–28 °C (64–82 °F)
	≤ 3 days

	Refrigerated
	2–8 °C (36–46 °F)
	≤ 1 week

	Frozen
	≤-18 °C (≤0 °F)
	Not recommended


 
Testing Procedure
Materials Provided
VITROS Chemistry Products ALT Slides
 
Materials Required but Not Provided
· VITROS Chemistry Products Calibrator Kit 3
· Quality control materials, such as VITROS Chemistry Products Performance Verifier I and II
· VITROS Chemistry Products 7% BSA
· VITROS Chemistry Products FS Diluent Pack 2 (BSA/Saline) (for on-analyzer dilution)
 
Operating Instructions
· Check reagent inventories at least daily to ensure that quantities are sufficient for the planned workload.
· For additional information, refer to the operating instructions for your system.
 
	IMPORTANT:
	Bring all fluids and samples to room temperature, 18–28 °C (64–82 °F), prior to analysis.


 
Sample Dilution
Serum and Plasma
If alanine aminotransferase activities exceed the system’s measuring (reportable or dynamic) range:
On-Analyzer Sample Dilution (VITROS 5600/XT 7600 Integrated and VITROS 250/350, 5,1 FS/4600/XT 3400 Chemistry Systems only)
Refer to the operating instructions for more information on the On-Analyzer Dilution Procedure. For VITROS 5600/XT 7600

Integrated and VITROS 5,1 FS/4600 Chemistry Systems, use VITROS Chemistry Products FS Diluent Pack 2 for the

dilution.

For VITROS 250/350 and VITROS XT 3400 Chemistry Systems, use VITROS 7% BSA for the dilution. 
Manual Sample Dilution
1. Dilute the sample with VITROS 7% BSA.
2. Reanalyze.
3. Multiply the results by the dilution factor to obtain an estimate of the original sample’s alanine aminotransferase activity.
 
 
Calibration
Required Calibrators
VITROS Chemistry Products Calibrator Kit 3
 
Calibrator Preparation, Handling, and Storage
Refer to the Instructions for Use for VITROS Calibrator Kit 3.
 
Calibration Procedure
Refer to the operating instructions for your system.
 
When to Calibrate
Calibrate:
· When the slide lot number changes.
· When critical system parts are replaced due to service or maintenance.
· When government regulations require.
For example, in the USA, CLIA regulations require calibration or calibration verification at least once every six months.
The VITROS ALT test may also need to be calibrated:
· If quality control results are consistently outside acceptable range.
· After certain service procedures have been performed.
For additional information, refer to the operating instructions for your system.
 
Calculations
Based on sequential readings of the slide’s reflectance at 340 nm over the defined incubation period, a rate of change in reflectance is determined. This rate is used in the software-resident multi-point rate calibration model to compute enzyme activity. Once a calibration has been performed for each slide lot, alanine aminotransferase activity in unknown samples can be determined from the rate of change in reflectance measured for each unknown test slide.
 
Validity of a Calibration
Calibration parameters are automatically assessed by the system against a set of quality parameters detailed in the

Coefficients and Limits screen on VITROS 250/350 Systems (on the VITROS 5600/XT 7600 Integrated and VITROS 5,1

FS/4600/XT 3400 Chemistry Systems, see the Review Assay Data screen). Failure to meet any of the pre-defined quality

parameters results in a failed calibration. The calibration report should be used in conjunction with quality control results to determine the validity of a calibration.
Measuring (Reportable or Dynamic) Range
 
	Conventional and SI Units (U/L)
	Alternate Units (µkat/L)

	6–1000
	0.10–16.70


For out-of-range samples, refer to “Sample Dilution.”
Traceability of Calibration
Values assigned to the VITROS Chemistry Products Calibrator Kit 3 for alanine aminotransferase are traceable to the alanine aminotransferase method recommended by the International Federation of Clinical Chemistry (IFCC), 8  adapted to a centrifugal analyzer at 37 °C.
 
Quality Control
Quality Control Material Selection
 
	IMPORTANT:
	VITROS Performance Verifiers are recommended for use with VITROS Chemistry and Integrated Systems. Evaluate the performance of other commercial control fluids for compatibility with this test before using for quality control.


· Control materials other than VITROS Performance Verifiers may show a difference when compared with other alanine aminotransferase methods if they:
–
Depart from a true human matrix.
–
Contain high concentrations of preservatives, stabilizers, or other nonphysiological additives.
· Enzyme activity might also vary with enzyme source, diluent temperature, and activation time during reconstitution.
· Do not use control materials stabilized with ethylene glycol.
Quality Control Procedure Recommendations
· Choose control levels that check the clinically relevant range.
· Analyze quality control materials in the same manner as patient samples, before or during patient sample processing.
· To verify system performance, analyze control materials:
–
After calibration.
–
According to local regulations or at least once each day that the test is being performed.
–
After specified service procedures are performed. Refer to the operating instructions for your system.
· If control results fall outside your acceptable range, investigate the cause before deciding whether to report patient results.
· For general quality control recommendations, refer to Statistical Quality Control for Quantitative Measurements: Principles and Definitions; Approved Guideline –Third Edition  9  or other published guidelines.
· For additional information, refer to the operating instructions for your system.
 
Quality Control Material Preparation, Handling, and Storage
Refer to the Instructions for Use for VITROS Chemistry Products Performance Verifier I and II or to other manufacturer's product literature.
 
Results
Reporting Units and Unit Conversion
The VITROS Chemistry and Integrated Systems may be programmed to report ALT results in conventional, SI, and alternate units.
 
	Conventional and SI Units
	Alternate Units

	U/L
	µkat/L (U/L x 0.0167)


 
Limitations of the Procedure
Known Interferences
None identified.
 
Other Limitations
Certain drugs and clinical conditions are known to alter alanine aminotransferase activity in vivo. For additional information, refer to one of the published summaries. 10  ,   11 
 
Expected Values
Reference Interval
These reference intervals are the central 95% of results from an internal study of 2444 apparently healthy adults (547 females and 1897 males).
 
	 
	Conventional and SI Units (U/L)
	Alternate Units (µkat/L)

	Adults
	13–69
	0.2–1.2

	Females
	9–52
	0.2–0.9

	Males
	21–72
	0.4–1.2


Each laboratory should confirm the validity of these intervals for the population it serves.
 
Performance Characteristics
Method Comparison
The plot and table show the results of a comparison of serum samples analyzed on the VITROS 750 System with those analyzed using the IFCC comparative method, 8  adapted to a centrifugal analyzer at 37 °C. Testing followed NCCLS Protocol EP9. 12 
The table also shows the results of comparisons of serum samples on the VITROS 250/350 and 950 Systems with the VITROS 750 System, and comparisons of the VITROS 5,1 FS System with the VITROS 950 System.
In addition, the table shows the results of a comparison with serum and plasma samples on the VITROS 5600 Integrated System and the VITROS 5,1 FS Chemistry System. The testing followed NCCLS Protocol EP9. 13 
Serum
 
	Conventional and SI Units
	Alternate Units

	VITROS 750 System (U/L)
	[image: image4.png]1200 -

0 200 400 600 800 1000 1200




	VITROS 750 System (µkat/L)
	[image: image5.png]4 6 8 10 12 14 16 18

o A
N A





	Comparative Method: Modified IFCC at 37 °C
	Comparative Method: Modified IFCC at 37 °C


 
	
	n
	Slope
	Correlation Coefficient
	Conventional and SI Units (U/L)
	Alternate Units (µkat/L)

	
	
	
	
	Range of Sample Activity
	Intercept
	Sy.x
	Range of Sample Activity
	Intercept
	Sy.x

	750 vs. comparative method
	193
	1.02
	0.997
	8–911
	+9.1
	20.4
	0.1–15.2
	+0.15
	0.34

	250 vs. 750
	60
	1.01
	0.999
	14–373
	-0.6
	2.6
	0.2–6.2
	-0.01
	0.04

	950 vs. 750
	122
	0.99
	0.999
	9–988
	-0.9
	2.4
	0.2–16.5
	-0.02
	0.04

	5,1 FS†  vs. 950
	134
	1.00
	1.000
	11–886
	+1.7
	2.3
	0.2–14.8
	+0.03
	0.04

	5600†† vs. 5,1 FS† 
	110
	1.00
	1.000
	9–954
	0.0
	3.1
	0.2–15.9
	0.00
	0.05


† Performance characteristics for the VITROS 5,1 FS System are applicable to the VITROS 4600 System.

†† Performance characteristics for the VITROS 5600 System are applicable to the VITROS XT 3400 and XT 7600 Systems. 
Precision
Precision was evaluated with quality control materials on VITROS 250/350 and 5,1 FS Systems following NCCLS Protocol EP5. 14  Precision was also evaluated with quality control materials on VITROS 5600 Integrated System following NCCLS protocol EP5. 15 
The data presented are a representation of test performance and are provided as a guideline. Variables such as sample handling and storage, reagent handling and storage, laboratory environment, and system maintenance can affect reproducibility of test results.
 
	 
	Conventional and SI Units (U/L)
	Alternate Units (µkat/L)
	Within Lab CV% ** 
	No. Observ.
	No. Days

	
	Mean Activity
	Within Day SD * 
	Within Lab SD ** 
	Mean Activity
	Within Day SD * 
	Within Lab SD ** 
	
	
	

	250
	37
	3.5
	4.1
	0.6
	0.06
	0.07
	11.2
	78
	20

	
	204
	3.8
	4.9
	3.4
	0.06
	0.08
	2.4
	78
	20

	5,1 FS† 
	44
	1.6
	2.9
	0.7
	0.03
	0.05
	6.6
	96
	22

	
	187
	1.9
	3.5
	3.1
	0.03
	0.06
	1.9
	96
	22

	5600††
	33
	1.3
	1.8
	0.6
	0.02
	0.03
	5.5
	88
	22

	
	170
	1.7
	3.1
	2.8
	0.03
	0.05
	1.8
	88
	22


* Within Day precision was determined using two runs/day with at least two replications.
** Within Lab precision was determined using a single lot of slides and calibrating weekly.
† Performance characteristics for the VITROS 5,1 FS System are applicable to the VITROS 4600 System.

†† Performance characteristics for the VITROS 5600 System are applicable to the VITROS XT 7600 System.
 
Specificity
Substances that Do Not Interfere
The substances listed in the table were tested with VITROS ALT Slides following NCCLS Protocol EP7 16 and found not to interfere, bias <8.3U/L, at the concentration shown.
 
	Compound
	Concentration
	
	Compound
	Concentration

	AST
	200 U/L
	200 U/L
	
	Glutathione
	1 mg/dL
	33 µmol/L

	Ascorbic Acid
	3 mg/dL
	170 µmol/L
	
	Intralipid
	800 mg/dL
	8 g/L

	Bilirubin
	40 mg/dL
	684 µmol/L
	
	Salicylic acid
	35 mg/dL
	2.5 mmol/L

	Ethanol
	300 mg/dL
	65 mmol/L
	
	
	
	


 
References
1. Tietz NW (ed). Fundamentals of Clinical Chemistry. ed. 3. Philadelphia: WB Saunders; 369–371; 1987.
2. CLSI.Protection of Laboratory Workers from Occupationally Acquired Infections; Approved Guideline – Fourth Edition. CLSI document M29-A4. Wayne, PA: Clinical and Laboratory Standards Institute; 2014.
3. Calam RR. Specimen Processing Separator Gels: An Update. J Clin Immunoassay. 11:86–90; 1988.
4. Young DS. Effects of Preanalytical Variables on Clinical Laboratory Tests. Washington D.C.: AACC Press; 3–7; 1993.
5. CLSI. Collection of Diagnostic Venous Blood Specimens. 7th ed. CLSI standard GP41. Wayne, PA: Clinical and Laboratory Standards Institute; 2017.

6. NCCLS. Procedures and Devices for the Collection of Diagnostic Capillary Blood Specimens; Approved Standard – Fifth Edition.NCCLS document H4-A5 [ISBN 1-56238-538-0]. CLSI, 940 West Valley Road, Suite 1400, Wayne, Pennsylvania 19087-1898 USA, 2004.
7. Clinical Laboratory Handbook for Patient Preparation and Specimen Handling. Fascicle VI: Chemistry/Clinical Microscopy. Northfield, IL: College of American Pathologists; 1992.
8. Bergmeyer HU, Horder M, Rej R. Approved Recommendation (1985) on IFCC Methods for the Measurement of Catalytic Concentration of Enzymes. Part 3. IFCC Method for Alanine Aminotransferase. J. Clin. Chem. Clin. Biochem. 24:481; 1986.
9. CLSI. Statistical Quality Control for Quantitative Measurements: Principles and Definitions; Approved Guideline – Third Edition. CLSI document C24-A3 (ISBN 1-56238-613-1). CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898 USA; 2006.
10. Young DS. Effects of Drugs on Clinical Laboratory Tests. ed. 4. Washington D.C.: AACC Press; 1995.
11. Friedman RB, Young DS. Effects of Disease on Clinical Laboratory Tests. Washington, D.C.: AACC Press; 1990.
12. CLSI. Method Comparison and Bias Estimation Using Patient Samples; Approved Guideline. CLSI Document EP9. CLSI, 940 West Valley Road, Suite 1400, Wayne, Pennsylvania 19087-1898 USA; 1995.
13. NCCLS. Method Comparison and Bias Estimation Using Patient Samples; Approved Guideline. NCCLS document EP9-A2 [ISBN 1-56238-472-4]. CLSI, 940 West Valley Road, Suite 1400, Wayne, Pennsylvania 19087-1898 USA; 2002.
14. CLSI. User Evaluation of Precision Performance with Clinical Chemistry Devices. CLSI Document EP5. CLSI, 940 West Valley Road, Suite 1400, Wayne, Pennsylvania 19087-1898 USA; 1992.
15. NCCLS. Evaluation of Precision Performance of Quantitative Measurement Methods; Approved Guideline – Second Edition. NCCLS document EP5-A2 [ISBN 1-56238-542-9]. CLSI, 940 West Valley Road, Suite 1400, Wayne, PA 19087-1898 USA; 2004
16. CLSI. Interference Testing in Clinical Chemistry. CLSI Document EP7. CLSI, 940 West Valley Road, Suite 1400, Wayne, Pennsylvania 19087-1898 USA; 1986.
 
Glossary of Symbols
[image: image6.png]The following symbols may have been used in the labeling of this product.

Do Not Reuse

Use by or Expiration Date
(Year-Month-Day)

el

LOT Batch Code or Lot Number

Serial Number

Catalog Number or
Product Code

m

RO [

Caution

Keep Dry (Protect from
Moisture/Humidity)

Manufacturer

Date of Manufacture

Authorized Representative
in the European Community

Corrosive

Health Hazards

/ﬂf Upper Limit of Temperature
Lower Limit of Temperature
Temperature Limitation

Consult Instructions for Use

Attention: The Instructions for
Use (IFU) has been updated

For use in Slide Supply 1
For use in Slide Supply 2

S| Units
Conventional Units

Value

Flammable

Acute Toxicity

® S [ilipgE=~

Range

Range of Means

Midpoint

Revised

Supersedes

Contains Sufficient for

n" Tests

in vitro Diagnostic

Medical Device

Der Griine Punkt (the Green Dot).
Manufacturer follows certain

packaging material waste disposal
management regulations

Estimated within-lab SD

Serious Health Hazards

Environmental or
Aquatic Toxicity





 
Revision History
 
	Date of Revision
	Version
	Description of Technical Changes

	 
	 
	 


 
	When this Procedure is replaced, sign and date below and retain as specified by local regulations or laboratory policies, as appropriate.

	 
	 
	 

	Signature
	 
	Obsolete Date


 
 
	  
	This Procedure was derived from the following Ortho Clinical Diagnostics document: Pub. No. MP2-36_EN Version 11.0
	The information in this Procedure document has been provided to customers using Ortho Clinical Diagnostics products and was current as of the publication date. This document is not intended for commercial sale or distribution and may not be provided to other users without the written permission of Ortho Clinical Diagnostics. The Instructions for Use documents are the authorized source of information for VITROS Chemistry Products and VITROS Immunodiagnostic Products.


	
	Ortho Clinical Diagnostics is not responsible for any editing, line alteration or other revision, intentional or otherwise, to this information.
	



	

	Ortho Clinical Diagnostics is not responsible for any editing, line alteration or other revision, intentional or otherwise, to this information.
	


	
	Ortho Clinical Diagnostics is not responsible for any editing, line alteration or other revision, intentional or otherwise, to this information.
	




[image: image7.png]